## Slide 1: Improving image quality of synthetically zoomed tomographic image
-Neha Dixit

## Slide 2: Presentation layout
Tomography basics/background
Problem statement and challenges faced
Existing approaches
Proposed approach
Results
Conclusion
Future work

## Slide 3: Tomographic image reconstruction
Tomography is a non-invasive imaging technique which helps to visualize the internal structures of an object
Cross-sectional images of body or an object are generated from sinogram data (radon transform).
Tomographic imaging modalities includes CT ,MRI, nuclear-medicine (SPECT and PET).

## Slide 4
X ray image
CT image
Source:Radiology picture of the day
Source : Dicts.info
Notes: Why tomography was introduced and how it is different from X-ray

## Slide 5
Concept of tomography

## Slide 6
Data Acquisition
CT scanner
PET scanner
Source: Jiri and Jan,
Medical image processing, reconstruction,
and restoration, concepts and methods.
Notes: Data acquisition in CT and PET will mention here that we will concentrate only on these two .

## Slide 7: Radon Transform
The Radon Transform (RT) is an integral transform whose inverse is used to reconstruct images
f(x,y) : a function defined on Cartesian coordinates (x,y)
p(r,φ): 1-D sinogram of f(x,y) defined over rotated (r,s) coordinate system

## Slide 8
=
=
Relationship between coordinates

## Slide 9: Radon transform…..
The RT of a distribution f(x, y) is given by
pφ(r)= p(r,φ)=R(f(x,y)
=
Writing the above equation in terms of (r,s)
pφ(r)= R(f(x,y)
=

## Slide 10
RT is nothing but the sinogram p(r,φ) of size 180x180
An Object and its Radon Transform

## Slide 11: Image Reconstruction
Given the Radon Transform of an object to reconstruct the cross-section image one of the following methods can be used:
Back projection
Filtered back projection
Iterative reconstruction

## Slide 12: Back Projection
The Back Projection (BP) operator is given as
b(x,y)=B(p(r,φ))= 0∫Π p(xcosΦ+ysinΦ, Φ) dΦ
Discrete version of BP is given as
B(p(nr,mφ))
= )+yjsin())
Drawback of direct BP
To overcome or minimize the above problem a well known algorithm Filtered Back Projection (FBP) is being used generally to reconstruct the images.

## Slide 13: Filtered Back Projection
Take 1-D Fourier transform of the sinogram data p(r,φ) for every view angle . Let this be denoted as P φ(ω).
Since convolution is equivalent to multiplication in Fourier domain, so 1-D Fourier transform is multiplied with 1-D ramp (|ω\) filter function to nullify the effect of PSF.
The filtered sinogram which is in the frequency domain is then transformed back to spatial domain by taking the 1-D inverse Fourier transform.
The final image is generated by backprojecting the processed sinogram.

## Slide 14
Filtered Back Projection
Cut of
frequency
Cut of frequency
Radon Transform
Back projection
1-D Fourier Transform
of sinogram
Back projection
Filtered Back Projection
(a)
(b)
(c)
(d)
(e)
(a) Shep-logan Phantom, (b) Sinogram , (c)Reconstructed image after simple back projection, (d) ramp filter, (e) Image reconstructed using FBP
Inverse Fourier
Transform of
filtered sinogram

## Slide 15: Iterative Reconstruction
x image matrix.
y sinogram data.
Drawbacks of MLEM
- Resulting image is noisy
- Slow convergence rate
- Computationally expensive

## Slide 16: Why Iterative Reconstruction
Source : Iterativereconstruction methods in X-ray CT
Marcel Beister, Daniel Kolditz, Willi A. Kalender

## Slide 17: Motivation
Improving resolution of tomographic images has been an active area of research in the past few years.
Nuclear imaging modalities such as PET and SPECT produce images of poor resolution (typically 128x128).
Higher resolution images are often required
    to view details
    for multimodal registration
Problem Statement
Generating higher resolution tomographic images of good quality from given scanner (sinogram) data

## Slide 18: Existing methods for improving resolution
Direct upsampling using a suitable kernel.
  Computationally inexpensive but image quality is compromised
Super resolution techniques :
(i) projection data is obtained after applying sub-pixel shifts and rotation to the object [Kennedy et al TMI 2006]
      Impacts dosage, scan time and data integrity
      Not feasible to implement with a patient
(ii) reconstruction by union of low resolution images on shifted lattices (denoted as USL) [Chang et al Med. Phys. 2008]

## Slide 19
Union of Shifted Lattices (USL)

## Slide 20
Required low resolution images (marked by stars) required for deriving an up- sampled (by 4) image using USL :
16 with SR algorithm [Chang et al Med. Phys. 2008]
7 with ISR-1 algorithm [Chang et al Med. Phys. 2009]
4 with ISR-2 algorithm [Chang et al Med. Phys. 2009]

## Slide 21: Proposed Solutions
An alternative lattice: In this approach, we use samples defined on single hexagonal lattice for generating upsampled images.
Multiple lattices: Here, an up-sampled image is generated using a super resolution technique by combining low resolution images which are defined on two rotated lattices (denotes as URL).
    Both hexagonal and square lattices were considered

## Slide 22
The Radon Transform (RT) is an integral transform whose inverse is used to reconstruct images
f(x,y) : a function defined on hexagonal coordinates (x,y)
p(r,φ): 1-D sinogram of f(x,y) defined over rotated (r,s) coordinate system
Radon transform for Hex Lattice

## Slide 23
Relationship between coordinates
=
=

## Slide 24: Radon transform for Hex Lattice
The Radon transform (RT) of a distribution f(x, y) is given by
pφ(r)=R(f(x,y))
=
Writing equation (2) in terms (1)
pφ(r)= R(f(x,y))
=

## Slide 25
Sinogram
(n* Φ)
DFT(1D)
Filtering
Inverse DFT
(1D)
Back Projection
(on discrete lattice)
Object
(continuous)
Reconstructed Image
2D Tomographic Reconstruction Using Filtered Back Projection

## Slide 26
Projection data
(sinogram) (n*Φ)
Filtered back projection on square lattice
Filtered back projection on hexagonal lattice
HK
Bicubic interpolation
Non-linear splines
SK
Up-sampling (by k) pipeline
Implementation details
Up-sampled (by k) image defined on square lattice:
, : upsampled images obtained with the square and hexagonal lattices

## Slide 27: Super Resolution (Multiple lattice)
Notes: How shall I approach this slide? The description is too big.

## Slide 28: Super resolution ….
, : sinograms acquired for the object without and with rotation.
, Noise
: BP operator that generates an image onto a lattice rotated by an angle .
+
+
+
Notes: Shall I include fourier domain equations also.

## Slide 29: Implementation details
Up-sampled image:+)⊗h
h: interpolation function
, : upsampled images obtained with the square and hexagonal lattices
Projection data
(sinogram) (n*Φ)
Filtered back projection on
lattice 1
Filtered back projection on lattice 2
Combination of samples for up-sampling by k
URLk
Pipeline for generating URL images, lattice 2 is a rotated version of lattice 1

## Slide 30: Implementation details …..
Reference images for Assessment of proposed method
Direct upsampled images Sk(ref1): The input sinogram was filtered and back projected onto a square lattice using the FBP algorithm. This was next up-sampled using bi-cubic interpolation to obtain image Sk.
Union of shifted lattices USLk (ref2): The FBP algorithm was used to generate sub-pixel shifted images defined on square lattices. [J.A.Fessler, image reconstruction toolbox],[Vandewalle , Super resolution toolbox ]
Projection data
(sinogram) (n*Φ)
Up-sampled SR image
Up-sampled image (bicubic interpolation)
Ref1:
USLk
Ref2:
Sk
Filter back projection on to shifted square lattice
Image with origin at (0,0)
Pipeline for Generating reference images for validation of results

## Slide 31: Phantoms
Concentric rings
Size:151x180
Lines
Size:121x180
Dots
Size:101x180
Hoffman brain
Size:396x315x91
NEMA
Size:293x280

## Slide 32: Quality Metric
Energy Distribution plots:
Energy Distribution plots in radial and
angular (wedge) directions [Bill,IQM toolbox]

## Slide 33: Quality Metric ……
Line profile:
CR=(/(
-m and n are number of peaks and troughs
-Pi is the amplitude of the ith peak
-Tj is the amplitude value of the j th trough.
Line profiles of the 1st row for Dots.

## Slide 34: Upsampled ROI(by 4) of Concentric Rings
S4
H4
USL4
URL4sq‏
URL4hex‏

## Slide 35: Spectra of upsampled concentric rings
S4
H4
USL4
URL4sq‏
URL4hex‏

## Slide 36: Wedge Plot -Concentric rings

## Slide 37: Radial Plot –Concentric rings

## Slide 38: Upsampled ROI(by4)- Line
S4
H4
USL4
URL4sq‏
URL4hex‏

## Slide 39: Upsampled ROI(by4)-Dots
S4
H4
USL4
URL4sq‏
URL4hex‏

## Slide 40: Intensity Profile of the penultimate line
 | S4 | H4 | USL4 | URL4sq | URL4hex
Contrast ratio | 0.5406 | 0.5262 | 0.5645 | 0.5717 | 0.5581

## Slide 41: Intensity Profile of the last line
 | S4 | H4 | USL4 | URL4sq | URL4hex
Contrast ratio | 0.6086 | 0.5874 | 0.6516 | 0.6638 | 0.6534

## Slide 42: Upsampled ROI(by4)-Dots(OSEM)
S6
USL6
URL6sq‏
 | S4 | USL4 | URL4sq
Contrast ratio | 0.6964 | 0.7761 | 0.7627

## Slide 43: Intensity Profile of the last line

## Slide 44: Nema- Upsampling factor 4
S4
H4
USL4
URL4sq‏
URL4hex‏

## Slide 45: Spectra of NEMA
S4
H4
USL4
URL4sq‏
URL4hex‏

## Slide 46
Line profiles of the smallest sphere
-NEMA phantom

## Slide 47: Radial Energy Plot –NEMA phantom

## Slide 48: Hoffman Brain Phantom (upsampling factor 4)
S4
USL4
URL4hex‏

## Slide 49: Hoffman Brain Phantom: Spectra
S4
USL4
URL4hex‏

## Slide 50: Line profile
Selected Line
Images | Difference in gray values of trough region |  |  |  |  |  |  |  |  |  |  | Avg
URL-S | -2 | -2 | -3 | -2 | -4 | -5 | -3 | -1 | -2 | -3 | -1 | -2.5
USL-S | -1 | -1 | -2 | 0 | -2 | -2 | -1 | +1 | 0 | -2 | -1 | 1.0
Images | Difference in gray values of crest region |  |  |  |  |  |  |  |  |  |  | Avg
URL-S | +5 | +7 | +7 | +7 | +7 | +6 | +4 | +5 | +5 | +5 | +5 | 5.7
USL-S | +2 | +3 | +3 | +3 | +4 | +4 | +3 | +3 | +3 | +3 | +2 | 3.0

## Slide 51: Hoffman Phantom: slice 19
S2
USL2
URL2hex‏
ROI
ROI : taken from the 19th slice of the Hoffman phantom
- used for computing noise spectra
Noise spectra of the ROI

## Slide 52: Summary
Problem addressed: Generating higher resolution tomographic images of good quality from given scanner (sinogram) data
Proposed Solutions:
  Use a Hexagonal lattice for reconstruction
  Use a combination of rotated lattices (hex or square) for reconstruction

## Slide 53: Conclusion
Single Lattice:
From the results it is evident that Hexagonal and Square lattices yield upsampled images that are of similar quality.
However, use of Hexagonal lattice incurs lesser computation cost, better representation of signal, additional equi-distance neighbor and better representation of curved structures.

## Slide 54: Conclusion ...
Multiple Lattice:
Images generated by URL scheme are of better quality when compared with direct upsampling method as well as of comparable quality to the image obtained by ISR-2 algorithm [Chang et al Med. Phys. 2009]
URL saves storage space and computational cost.
For example for generating K times upsampled image, K2 shifted images are needed. ISR-1 and ISR-2 require 2*(k-1) and K images respectively. Our proposed method requires 2 images for any upsampling (k=2…6)factor.

## Slide 55: Future Work
Explore different interpolation kernels: hex splines etc.
Explore design of adaptive kernel.
Multiple lattices: Optimal rotational angle could be computed theoretically.
Explore GPU platform for further speedup.

## Slide 56
Reference:
[1] J. A. Kennedy, O. Israel, A. Frenkel, R. Bar-Shalom, and H. Azhari,
“Super resolution in PET imaging”, TMI 25(2), 2006.
[2] G. Chang, T. Pan, J. Clark Jr., O.R. Mawlawi, “Comparison between two
super resolution implementations in PET imaging” Med. Phys, 2009.
[3] G Chang, T Pan, J Clark, Jr, O.R. Mawlawi, “Optimization of super-
resolution processing using incomplete image sets” Med. Phys, 2008.
[4] P.Vandewalle, K. Krichane, P. Zbinden, Super resolution toolbox
(http://lcavwww.epfl.ch/software)
[5] J.A.Fessler, image reconstruction toolbox
(http://www.eecs.umich.edu/fessler/code)
[6] N.B. Nill and B.H. Bouzas, "Objective Image Quality Measure Derived
from Digital Image Power Spectra”, Optical Engineering, vol.31, 1992

## Slide 57: Publication
N. Dixit, N. V. K. Medathati and J. Sivaswamy, "Synthetic Zooming of Tomographic Images by Combination of Lattices", Nuclear Science Symposium Medical Imaging Conference, Orlando, 2009.
N. Dixit and J. Sivaswamy, "Novel Approach to Generate Up-sampled Tomographic Images using Combination of Rotated Hexagonal Lattices ", National Conference on Communication, Chennai, 2010.

## Slide 58: Super resolution concept
Fourier transform of function f(x)
Introducing a small shift in x direction
If = 0.5 (sub-pixel shift)
() =

## Slide 59: Super resolution: Hex lattice

## Slide 60: Hounsfield Table
Substance | HU
Air | −1000
Lung | −700
Soft Tissue | −300 to -100
Fat | −84
Water | 0
CSF | 15
Blood | +30 to +45
Muscle | +40
Bone | +700 to +3000 (dense bone)